- Rules and Tips for the Microscope Operation

- Turn on the Microscope (following the orders):

  - Arc Lamp Power.

  - Shutter/Filter Wheels Controller.

  - Halogen Shutter (Uniblitz, on the backside).

  - Cascade Camera (on the backside).

  - Zeiss Microscope (on the right side).

  - Computer

- Shut off the Microscope (following the orders):

  - lower the objective, rotate the objective inside

  - take out the samples

  - raise the objective again

  - Clean the Oil Objectives (if any).

  - Make sure the Objective is in Up position.

  - turn off Computer.

  - turn off Zeiss Microscope (on the right side).

  - turn off Cascade Camera (on the backside).

  - turn off Halogen Shutter (Uniblitz, on the backside).

  - turn off Shutter/Filter Wheels Controller.

  - turn off Arc Lamp Power

- Others:

- If you used oil objectives, please make sure to clean it.

  - cotton swabs + 70% Alcohol

  - Lens Papers + Lens Cleaning Solution

  - Dry Lens Papers

- Kohler Alignment

  - Focus on the specimen using low magnification

  - Close Field Diaphragm

  - Adjust Condenser up and down until a sharp hexagon is observed

  - Center the hexagon if needed with the two knobs

  - open the diaphragm until the hexagon touches the field edge simultaneously.

- Components to adjust for DIC images.

  - Polarizer (second ring filter on top)

  - DIC Prism in the Condenser (e.g. 100X Objective, should choose DIC III on LCD display)

  - Turn the prism under Objectives (Slide in to be functional).

  - Polarizer in the Motorized Diachroic (Analyzer, DIC).

- To get a focus, use 10x objective first, focus, hit the “objectives” button to switch the objective to higher magnifications, fine adjust the focus to get a better view.

- If using low voltage/intensity for Halogen Lamp and seeing yellow color, put the First ring filter in position to balance the color.